## Test Results
gVent Testing
black features are those required for non-ICU ventilators: https://buyandsell.gc.ca/sites/buyandsell.gc.ca/files/non-icu-ventilator-requirements_0.pdf
blue features are those speficied for ICU ventilators: https://buyandsell.gc.ca/sites/buyandsell.gc.ca/files/ventilatorspecs2020covid19outbreak.pdf - Refer to this document for comprehensive list, not all criteria included here
Test Date | 2020-03-31 | 2020-04-09
Prototype | 3 | 4
| no P/flow sensors | P/flow sensors inclusive
Capabilities re: Ventilator Modes
Controlled ventilation with:
pressure control | yes
volume control | -
Pressure Support | -
Synchronized intermittent mechanical ventilation (SIMV) - essentially PC or VC + PS | -
Back up ventilation 2º to low MV | -
Ventilator Parameters
FiO2 21-100% | -
PEEP range (Goal 0-30; 35cm H2O) | 0-25
Pressure support beyond PEEP (Goal of 60cm H2O) | -
Maximum Flowrate (goal 100; 150 SLP) | 56.05
Max Tidal Volume (goal of 1500ml;2000ml) | 1.955
Maximum Resp Rate (goal 45 bpm;150bpm) | 60
I:E Ratio Range (goal 1:1 - 1:4) | 1:1 - 1:5 | 1:1 - 1:5
Peak Inspiratory Pressure (80cm H2O) | 40
Minimum Inspiratory Pressure | 16
Inspiratory Time (Goal 0.1-5s) | -
Inspiratory rise time (0-.4s) | -
Flow trigger sensitivity (tested threshold) | -
Prssure trigger sensitivity (tested threshold) | -
Alarms (45-85dBA)
Peak Pressure
High | -
Low | -
Tidal Volume (optional for non-ICU)
High | -
Low | -
Respiratory Rate (optional for non-ICU)
High | -
Low | -
Minute Ventilation (TV*RR)
High | -
Low | -
FiO2
High | -
Low | -
PEEP
High | -
Low | -
Apnea alarm with back up rate & TV | -
Disconnection alarm
Extra considerations
Safety Relief Valve | Pressure limited at 40cmH20
Plateau Pressure Readings
Ability to connect various vent circuits?
Dispay/touch screen (all only for ICU vent)
Vent mode & breath type
Delivered FiO2
Airway pressures:
Peak
Mean
Plateau
PEEP
Respiratory Rate
TV
MV
I:E Ratio
Graphing of:
Pressure vs t
Flow vs. t
Volume vs. t
Flow/volume loops

## Sheet2
https://www.gov.uk/government/publications/coronavirus-covid-19-ventilator-supply-specification/rapidly-manufactured-ventilator-system-specification#ventilation
Ventilation
At least 1, optionally 2 modes of ventilation:
must have mandatory ventilation (for the deeply sedated and paralysed). The user can set a tidal volume and the output is a pressure regulated flow to achieve this volume, for example, pressure regulated volume control (PRVC), SIMV-PC
optional pressure support mode for those patients breathing to some extent themselves, for example, BIPAP. The user sets an inspiratory pressure and an expiratory pressure. The ventilator can sense when a patient starts to breathe in and apply the inspiratory pressure, then sense when the patient starts to breathe out and apply the expiratory pressure (this pressure is still positive but lower than the inspiratory pressure)
If the patient stops breathing in pressure support mode, it must failsafe automatically onto mandatory ventilation.
Inspiratory airway pressure, the higher pressure setting that is applied to make the patient breathe in:
plateau pressure should adapt to achieve volume and be limited to 35 cmH2O
peak pressure should be no more than 2 cmH2O greater than plateau pressure
ideally there should be a mechanical failsafe valve that opens at 40 cmH2O
Positive End Expiratory Pressure PEEP (usually called EPAP during pressure support mode). The lower pressure applied to the patients airway to allow them to breathe out, but not too much:
range 5 to 25 cm H2O adjustable in 5 cmH2O increment
patient breathing system must remain pressurised to at least the PEEP level setting at all times
Inspiratory:Expiratory ratio (I:E) (note, confusingly, it is actually E/I time). The proportion of each breathing cycle that is spent breathing in compared to breathing out:
2.0 (i.e. expiration lasts twice as long as inspiration)
optionally adjustable in the range 1.0 to 3.0
Respiratory Rate. The number of breathing cycles every minute:
range 10 to 30 breaths per minute in increments of 2 (only in mandatory mode) can be set by the user
Tidal Volume (Vt). The volume of gas flowing into the lungs during one inspiratory cycle:
must have at least one setting of 400ml +/- 10 ml
ideally 350ml and 450 ml options
optionally Range 250 to 600 ml in steps of 50ml
even more optionally up to 800 ml
optionally the ability to input body weight and have volume calculated as, for example, 6ml/kg of ideal body weight